## States
States | Substates
Rinse | Rinse
Preparation | Preparation | Dializer priming
Disinfection | Hot citric disinfection | Bleach disinfection | Hot rinse
Dialysis | Normal dialysis | ISO UF | Heparin-less dialysis
Self test | Self test

## Rinse
Alarms / Reasons | Reason 1 | Reason 2 | Reason 3 | Reason 4 | Reason 5 | Reason 6 | Reason 7 | Reason 8 | Reason 9 | Reason 10
Main power supply failure | EMI filter malfunction | Battery PCB connector issues | Battery PCB issue
Red probe Open | Microswitch issues | Connector issue
Blue probe open | Microswitch issues | Connector issue
Blue shunt connector Open | Microswitch issues | Connector issue | Connector in Hydromatrix PCB issue | Hydromatrix PCB issue | Main CPU issue
Red shunt connector Open | Microswitch issues | Connector issue | Connector in Hydromatrix PCB issue | Hydromatrix PCB issue | Main CPU issue
Water spillage on the floor | Pipe disconnection
UF pump under run | piston pump feedback sensor issue | Pump stuck | Drive Malfunctioning
UF pump Over run | piston pump feedback sensor issue | Drive Malfunctioning
No Flow | Flow sensor issue | RO inlet stud broken | PRV malfunction | V1 malfunction | Flow pump malfunction | LSC malfunction
Low Flow | Flow sensor issue | RO inlet stud broken | PRV malfunction | V1 malfunction | Flow pump malfunction | LSC malfunction
Acid pump overrun | piston pump feedback sensor issue |  | Drive Malfunctioning
Acid pump underrun | piston pump feedback sensor issue |  | Pump stuck | Drive Malfunctioning
Bicarb pump overrun | piston pump feedback sensor issue |  | Drive Malfunctioning
Bicarb pump underrun | piston pump feedback sensor issue |  | Pump stuck | Drive Malfunctioning
High dialyzer inlet pressure | NA
High dialyzer outlet pressure | NA
Heater Temperature High | NA
Communication lost | USB cable issue | Power cord issue
Low Arterial pressure | NA
High Arterial pressure | NA
Low venous pressure | NA
High vinous pressure | NA
Blood pump door open | NA
Blood pump stopped | NA
Blood not detected | NA
Air bubble detected | NA
Blood leak detected | NA
Low dialysate temperature | NA
High dialysate temperature | NA
Low dialysate conductivity | NA
High dialysate conductivity | NA
Heparin pump stopped | NA
High RO inlet temperature | NA
BP underrun | NA
BP overrun | NA
Heparin underrun | NA
Heparin overrun | NA
Low RO inlet temperature | NA
Issues
Air bubbles in probe tubes
Leakage in CSA stud | Stud breakage

## Preparation
Alarms / Reasons | Reason 1 | Reason 2 | Reason 3 | Reason 4 | Reason 5 | Reason 6 | Reason 7 | Reason 8 | Reason 9 | Reason 10
Main power supply failure | EMI filter malfunction | Battery PCB connector issues | Battery PCB issue
Red probe Open | Microswitch issues | Connector issue
Blue probe open | Microswitch issues | Connector issue
Blue shunt connector Open | Microswitch issues | Connector issue | Connector in Hydromatrix PCB issue | Hydromatrix PCB issue | Main CPU issue
Red shunt connector Open | Microswitch issues | Connector issue | Connector in Hydromatrix PCB issue | Hydromatrix PCB issue | Main CPU issue
UF pump under run | piston pump feedback sensor issue | Pump stuck | Drive Malfunctioning
UF pump Over run | piston pump feedback sensor issue | Drive Malfunctioning
No Flow | Flow sensor issue | RO inlet stud broken | PRV malfunction | V1 malfunction | Flow pump malfunction | LSC malfunction
Low Flow | Flow sensor issue | RO inlet stud broken | PRV malfunction | V1 malfunction | Flow pump malfunction | LSC malfunction
High Flow | Flow sensor issue
Acid pump overrun | piston pump feedback sensor issue |  | Drive Malfunctioning
Acid pump underrun | piston pump feedback sensor issue |  | Pump stuck | Drive Malfunctioning | Pump suction issue | Bicarb pump issue
Bicarb pump overrun | piston pump feedback sensor issue |  | Drive Malfunctioning
Bicarb pump underrun | piston pump feedback sensor issue |  | Pump stuck | Drive Malfunctioning
High dialyzer inlet pressure | Pressure sensor issue | Flow pump issue
High dialyzer outlet pressure | Pressure sensor issue | Flow pump issue | CV1 issue | V18 issue | UF filter blockage
Heater Temperature High | Heater PCB connector issue | Heater PCB issue | TS2 sensor issue
Communication lost | USB cable issue | Power cord issue
Low Arterial pressure | NA
High Arterial pressure | NA
Low venous pressure | NA
High vinous pressure | NA
Blood pump door open | NA
Blood pump stopped | NA
Blood not detected | NA
Air bubble detected | NA
Blood leak detected | NA
Low dialysate temperature
High dialysate temperature
Low dialysate conductivity
High dialysate conductivity
Heparin pump stopped | NA
High RO inlet temperature
BP underrun
BP overrun
Heparin underrun
Heparin overrun
Low RO inlet temperature
Issues
Blood pump started running automatically
Not observed pop up message for saline prime | Conductivity or Temperature not achieved for set parameter
Bicarb shunt connector issue
Blood pump speed not changing
Blood pump not starting after pressing priming button during preparation
Leakage in Red shunt coupler
Heparin pump got stuck | Proper fix of Back cover
Extracorporeal blood pump control key's LED and key's are not working | Extracorporeal keys replacement
Conductivity Fluctuation between CS2 and CS3

## Self test
Alarms / Reasons | Reason 1 | Reason 2 | Reason 3 | Reason 4 | Reason 5 | Reason 6 | Reason 7 | Reason 8 | Reason 9 | Reason 10
Main power supply failure | EMI filter malfunction | Battery PCB connector issues | Battery PCB issue
Red probe Open | Microswitch issues | Connector issue
Blue probe open | Microswitch issues | Connector issue
Blue shunt connector Open | Microswitch issues | Connector issue | Connector in Hydromatrix PCB issue | Hydromatrix PCB issue | Main CPU issue
Red shunt connector Open | Microswitch issues | Connector issue | Connector in Hydromatrix PCB issue | Hydromatrix PCB issue | Main CPU issue
UF pump under run | piston pump feedback sensor issue | Pump stuck | Drive Malfunctioning
UF pump Over run | piston pump feedback sensor issue | Drive Malfunctioning
No Flow | Flow sensor issue | RO inlet stud broken | PRV malfunction | V1 malfunction | Flow pump malfunction | LSC malfunction
Low Flow | Flow sensor issue | RO inlet stud broken | PRV malfunction | V1 malfunction | Flow pump malfunction | LSC malfunction
High Flow | Flow sensor issue
Acid pump overrun | piston pump feedback sensor issue |  | Drive Malfunctioning
Acid pump underrun | piston pump feedback sensor issue |  | Pump stuck | Drive Malfunctioning
Bicarb pump overrun | piston pump feedback sensor issue |  | Drive Malfunctioning
Bicarb pump underrun | piston pump feedback sensor issue |  | Pump stuck | Drive Malfunctioning
High dialyzer inlet pressure | Pressure sensor issue | Flow pump issue
High dialyzer outlet pressure | Pressure sensor issue | Flow pump issue | CV1 issue | V18 issue | UF filter blockage
Heater Temperature High | Heater PCB connector issue | Heater PCB issue | TS2 sensor issue
Communication lost | USB cable issue | Power cord issue
Low Arterial pressure
High Arterial pressure
Low venous pressure
High vinous pressure
Blood pump door open
Blood pump stopped
Blood not detected
Air bubble detected
Blood leak detected | BLD sensor failed
Low dialysate temperature
High dialysate temperature
Low dialysate conductivity
High dialysate conductivity
Heparin pump stopped
High RO inlet temperature
BP underrun
BP overrun
Heparin underrun
Heparin overrun
Low RO inlet temperature
Venous clamp | Venous clamp failed
Issues
BLD failed | Air bubble
Venous clamp getting ON and OFF after self-test completed
Heparin pump failed

## Dializer priming
Alarms / Reasons | Reason 1 | Reason 2 | Reason 3 | Reason 4 | Reason 5 | Reason 6 | Reason 7 | Reason 8 | Reason 9 | Reason 10
Main power supply failure | EMI filter malfunction | Battery PCB connector issues | Battery PCB issue
Red probe Open | Microswitch issues | Connector issue
Blue probe open | Microswitch issues | Connector issue
Blue shunt connector Open | Microswitch issues | Connector issue | Connector in Hydromatrix PCB issue | Hydromatrix PCB issue | Main CPU issue
Red shunt connector Open | Microswitch issues | Connector issue | Connector in Hydromatrix PCB issue | Hydromatrix PCB issue | Main CPU issue
UF pump under run | piston pump feedback sensor issue | Pump stuck | Drive Malfunctioning
UF pump Over run | piston pump feedback sensor issue | Drive Malfunctioning
No Flow | Flow sensor issue | RO inlet stud broken | PRV malfunction | V1 malfunction | Flow pump malfunction | LSC malfunction
Low Flow | Flow sensor issue | RO inlet stud broken | PRV malfunction | V1 malfunction | Flow pump malfunction | LSC malfunction
High Flow | Flow sensor issue
Acid pump overrun | piston pump feedback sensor issue |  | Drive Malfunctioning
Acid pump underrun | piston pump feedback sensor issue |  | Pump stuck | Drive Malfunctioning
Bicarb pump overrun | piston pump feedback sensor issue |  | Drive Malfunctioning
Bicarb pump underrun | piston pump feedback sensor issue |  | Pump stuck | Drive Malfunctioning
High dialyzer inlet pressure | Pressure sensor issue | Flow pump issue
High dialyzer outlet pressure | Pressure sensor issue | Flow pump issue | CV1 issue | V18 issue | UF filter blockage
Heater Temperature High | Heater PCB connector issue | Heater PCB issue | TS2 sensor issue
Communication lost | USB cable issue | Power cord issue
Low Arterial pressure
High Arterial pressure
Low venous pressure
High venous pressure
Blood pump door open | alarm will not generate for this
Blood pump stopped | alarm will not generate for this
Blood not detected | alarm will not generate for this
Air bubble detected
Blood leak detected
Low dialysate temperature
High dialysate temperature
Low dialysate conductivity
High dialysate conductivity
Heparin pump stopped
High RO inlet temperature
BP underrun
BP overrun
Heparin underrun
Heparin overrun
Low RO inlet temperature
Issues
Priming completed without blood pump rotation
Dripchambers filling up completely | Luer connector issue

## Hot citric disinfection
Alarms / Reasons | Reason 1 | Reason 2 | Reason 3 | Reason 4 | Reason 5 | Reason 6 | Reason 7 | Reason 8 | Reason 9 | Reason 10
Main power supply failure | EMI filter malfunction | Battery PCB connector issues | Battery PCB issue
Red probe Open | Microswitch issues | Connector issue
Blue probe open | Microswitch issues | Connector issue
Blue shunt connector Open | Microswitch issues | Connector issue | Connector in Hydromatrix PCB issue | Hydromatrix PCB issue | Main CPU issue
Red shunt connector Open | Microswitch issues | Connector issue | Connector in Hydromatrix PCB issue | Hydromatrix PCB issue | Main CPU issue
UF pump under run | piston pump feedback sensor issue | Pump stuck | Drive Malfunctioning
UF pump Over run | piston pump feedback sensor issue | Drive Malfunctioning
No Flow | Flow sensor issue | RO inlet stud broken | PRV malfunction | V1 malfunction | Flow pump malfunction | LSC malfunction | TS2 sensor was not proper
Low Flow | Flow sensor issue | RO inlet stud broken | PRV malfunction | V1 malfunction | Flow pump malfunction | LSC malfunction | TS2 sensor was not proper
High Flow | Flow sensor issue
Acid pump overrun | piston pump feedback sensor issue |  | Drive Malfunctioning
Acid pump underrun | piston pump feedback sensor issue |  | Pump stuck | Drive Malfunctioning
Bicarb pump overrun | piston pump feedback sensor issue |  | Drive Malfunctioning
Bicarb pump underrun | piston pump feedback sensor issue |  | Pump stuck | Drive Malfunctioning
High dialyzer inlet pressure | Pressure sensor issue | Flow pump issue
High dialyzer outlet pressure | Pressure sensor issue | Flow pump issue | CV1 issue | V18 issue | UF filter blockage
Heater Temperature High | Heater PCB connector issue | Heater PCB issue | TS2 sensor issue
Communication lost | USB cable issue | Power cord issue
Container empty
Low Arterial pressure | NA
High Arterial pressure | NA
Low venous pressure | NA
High vinous pressure | NA
Blood pump door open | NA
Blood pump stopped | NA
Blood not detected | NA
Air bubble detected | NA
Blood leak detected | NA
Low dialysate temperature | NA
High dialysate temperature | NA
Low dialysate conductivity | NA
High dialysate conductivity | NA
Heparin pump stopped | NA
High RO inlet temperature
BP underrun
BP overrun
Heparin underrun
Heparin overrun
Low RO inlet temperature
Issues
Water/hot steam/vapour in LSC tube | Orifice issue | TS2 faulty
Disinfection is not happening in Bicarb probe tube
Hot water not reaching everywhere | Bicarb pump issue
Container empty alarm not coming
Higher heater duty

## Bleach disinfection
Alarms / Reasons | Reason 1 | Reason 2 | Reason 3 | Reason 4 | Reason 5 | Reason 6 | Reason 7 | Reason 8 | Reason 9 | Reason 10
Main power supply failure | EMI filter malfunction | Battery PCB connector issues | Battery PCB issue
Red probe Open | Microswitch issues | Connector issue
Blue probe open | Microswitch issues | Connector issue
Blue shunt connector Open | Microswitch issues | Connector issue | Connector in Hydromatrix PCB issue | Hydromatrix PCB issue | Main CPU issue
Red shunt connector Open | Microswitch issues | Connector issue | Connector in Hydromatrix PCB issue | Hydromatrix PCB issue | Main CPU issue
UF pump under run | piston pump feedback sensor issue | Pump stuck | Drive Malfunctioning
UF pump Over run | piston pump feedback sensor issue | Drive Malfunctioning
No Flow | Flow sensor issue | RO inlet stud broken | PRV malfunction | V1 malfunction | Flow pump malfunction | LSC malfunction
Low Flow | Flow sensor issue | RO inlet stud broken | PRV malfunction | V1 malfunction | Flow pump malfunction | LSC malfunction
High Flow | Flow sensor issue
Acid pump overrun | piston pump feedback sensor issue |  | Drive Malfunctioning
Acid pump underrun | piston pump feedback sensor issue |  | Pump stuck | Drive Malfunctioning
Bicarb pump overrun | piston pump feedback sensor issue |  | Drive Malfunctioning
Bicarb pump underrun | piston pump feedback sensor issue |  | Pump stuck | Drive Malfunctioning
High dialyzer inlet pressure | Pressure sensor issue | Flow pump issue
High dialyzer outlet pressure | Pressure sensor issue | Flow pump issue | CV1 issue | V18 issue | UF filter blockage
Heater Temperature High | Heater PCB connector issue | Heater PCB issue | TS2 sensor issue
Communication lost | USB cable issue | Power cord issue
Container empty | NA
Low Arterial pressure | NA
High Arterial pressure | NA
Low venous pressure | NA
High vinous pressure | NA
Blood pump door open | NA
Blood pump stopped | NA
Blood not detected | NA
Air bubble detected | NA
Blood leak detected | NA
Low dialysate temperature | NA
High dialysate temperature | NA
Low dialysate conductivity | NA
High dialysate conductivity | NA
Heparin pump stopped | NA
High RO inlet temperature
BP underrun
BP overrun
Heparin underrun
Heparin overrun
Low RO inlet temperature
Issues
Disinfection is not happening in Bicarb probe tube

## Hot rinse
Alarms / Reasons | Reason 1 | Reason 2 | Reason 3 | Reason 4 | Reason 5 | Reason 6 | Reason 7 | Reason 8 | Reason 9 | Reason 10
Main power supply failure | EMI filter malfunction | Battery PCB connector issues | Battery PCB issue
Red probe Open | Microswitch issues | Connector issue
Blue probe open | Microswitch issues | Connector issue
Blue shunt connector Open | Microswitch issues | Connector issue | Connector in Hydromatrix PCB issue | Hydromatrix PCB issue | Main CPU issue
Red shunt connector Open | Microswitch issues | Connector issue | Connector in Hydromatrix PCB issue | Hydromatrix PCB issue | Main CPU issue
UF pump under run | piston pump feedback sensor issue | Pump stuck | Drive Malfunctioning
UF pump Over run | piston pump feedback sensor issue | Drive Malfunctioning
No Flow | Flow sensor issue | RO inlet stud broken | PRV malfunction | V1 malfunction | Flow pump malfunction | LSC malfunction | orifice issue
Low Flow | Flow sensor issue | RO inlet stud broken | PRV malfunction | V1 malfunction | Flow pump malfunction | LSC malfunction
High Flow | Flow sensor issue
Acid pump overrun | piston pump feedback sensor issue |  | Drive Malfunctioning
Acid pump underrun | piston pump feedback sensor issue |  | Pump stuck | Drive Malfunctioning
Bicarb pump overrun | piston pump feedback sensor issue |  | Drive Malfunctioning
Bicarb pump underrun | piston pump feedback sensor issue |  | Pump stuck | Drive Malfunctioning
High dialyzer inlet pressure | Pressure sensor issue | Flow pump issue
High dialyzer outlet pressure | Pressure sensor issue | Flow pump issue | CV1 issue | V18 issue | UF filter blockage
Heater Temperature High | NA
Communication lost | USB cable issue | Power cord issue
Low Arterial pressure | NA
High Arterial pressure | NA
Low venous pressure | NA
High vinous pressure | NA
Blood pump door open | NA
Blood pump stopped | NA
Blood not detected | NA
Air bubble detected | NA
Blood leak detected | NA
Low dialysate temperature | NA
High dialysate temperature | NA
Low dialysate conductivity | NA
High dialysate conductivity | NA
Heparin pump stopped | NA
High RO inlet temperature
BP underrun
BP overrun
Heparin underrun
Heparin overrun
Low RO inlet temperature
Issues
Low flow during post rinse
Hot water is not passing in red probe side | Acid pump getting stuck
Water/hot steam/vapour in LSC tube | Orifice issue | TS2 Malfunction
Hot water not reaching everywhere | Bicarb pump issue

## Normal dialysis
Alarms / Reasons | Reason 1 | Reason 2 | Reason 3 | Reason 4 | Reason 5 | Reason 6 | Reason 7 | Reason 8 | Reason 9 | Reason 10
Main power supply failure | EMI filter malfunction | Battery PCB connector issues | Battery PCB issue
Red probe Open | Microswitch issues | Connector issue
Blue probe open | Microswitch issues | Connector issue
Blue shunt connector Open | Microswitch issues | Connector issue | Connector in Hydromatrix PCB issue | Hydromatrix PCB issue | Main CPU issue
Red shunt connector Open | Microswitch issues | Connector issue | Connector in Hydromatrix PCB issue | Hydromatrix PCB issue | Main CPU issue
UF pump under run | piston pump feedback sensor issue | Pump stuck | Drive Malfunctioning
UF pump Over run | piston pump feedback sensor issue | Drive Malfunctioning
No Flow | Flow sensor issue | RO inlet stud broken | PRV malfunction | V1 malfunction | Flow pump malfunction | LSC malfunction
Low Flow | Flow sensor issue | RO inlet stud broken | PRV malfunction | V1 malfunction | Flow pump malfunction | LSC malfunction
High Flow | Flow sensor issue
Acid pump overrun | piston pump feedback sensor issue |  | Drive Malfunctioning
Acid pump underrun | piston pump feedback sensor issue |  | Pump stuck | Drive Malfunctioning
Bicarb pump overrun | piston pump feedback sensor issue |  | Drive Malfunctioning
Bicarb pump underrun | piston pump feedback sensor issue |  | Pump stuck | Drive Malfunctioning
High dialyzer inlet pressure | Pressure sensor issue | Flow pump issue
High dialyzer outlet pressure | Pressure sensor issue | Flow pump issue | CV1 issue | V18 issue | UF filter blockage
Heater Temperature High | Heater PCB connector issue | Heater PCB issue | TS2 sensor issue
Communication lost | USB cable issue | Power cord issue
Low Arterial pressure
High Arterial pressure | Venous clamp door issue
Low venous pressure
High venous pressure | Venous clamp door issue
Blood pump door open
Blood pump stopped
Blood not detected
Air bubble detected | Air bubbles | Tube not inserted properly
Blood leak detected
Low dialysate temperature
High dialysate temperature
Low dialysate conductivity
High dialysate conductivity
Heparin pump stopped
High RO inlet temperature
BP underrun
BP overrun
Heparin underrun
Heparin overrun
Low RO inlet temperature
Issues
Blood pump speed not changing
AP and VP got Stuck | Hemo cpu issue - reset
Leakage in Red shunt coupler
Air bubble alarm not working | ABD sensor issue
BLD failed to detect blood leak at dialyzer outlet
Jerking in the arterial drip chamber

## ISO UF
Alarms / Reasons | Reason 1 | Reason 2 | Reason 3 | Reason 4 | Reason 5 | Reason 6 | Reason 7 | Reason 8 | Reason 9 | Reason 10
Main power supply failure | EMI filter malfunction | Battery PCB connector issues | Battery PCB issue
Red probe Open | Microswitch issues | Connector issue
Blue probe open | Microswitch issues | Connector issue
Blue shunt connector Open | Microswitch issues | Connector issue | Connector in Hydromatrix PCB issue | Hydromatrix PCB issue | Main CPU issue
Red shunt connector Open | Microswitch issues | Connector issue | Connector in Hydromatrix PCB issue | Hydromatrix PCB issue | Main CPU issue
UF pump under run | piston pump feedback sensor issue | Pump stuck | Drive Malfunctioning
UF pump Over run | piston pump feedback sensor issue | Drive Malfunctioning
No Flow | Flow sensor issue | RO inlet stud broken | PRV malfunction | V1 malfunction | Flow pump malfunction | LSC malfunction
Low Flow | Flow sensor issue | RO inlet stud broken | PRV malfunction | V1 malfunction | Flow pump malfunction | LSC malfunction
High Flow | Flow sensor issue
Acid pump overrun | piston pump feedback sensor issue |  | Drive Malfunctioning
Acid pump underrun | piston pump feedback sensor issue |  | Pump stuck | Drive Malfunctioning
Bicarb pump overrun | piston pump feedback sensor issue |  | Drive Malfunctioning
Bicarb pump underrun | piston pump feedback sensor issue |  | Pump stuck | Drive Malfunctioning
High dialyzer inlet pressure | Pressure sensor issue | Flow pump issue
High dialyzer outlet pressure | Pressure sensor issue | Flow pump issue | CV1 issue | V18 issue | UF filter blockage
Heater Temperature High | Heater PCB connector issue | Heater PCB issue | TS2 sensor issue
Communication lost | USB cable issue | Power cord issue
Low Arterial pressure
High Arterial pressure
Low venous pressure
High vinous pressure
Blood pump door open
Blood pump stopped
Blood not detected
Air bubble detected
Blood leak detected
Low dialysate temperature
High dialysate temperature
Low dialysate conductivity
High dialysate conductivity
Heparin pump stopped
High RO inlet temperature
BP underrun
BP overrun
Heparin underrun
Heparin overrun
Low RO inlet temperature
Issues
Air bubble alarm not coming

## Heparin-less dialysis
Alarms / Reasons | Reason 1 | Reason 2 | Reason 3 | Reason 4 | Reason 5 | Reason 6 | Reason 7 | Reason 8 | Reason 9 | Reason 10
Main power supply failure
Red probe Open
Blue probe open
Blue shunt connector Open
Red shunt connector Open
UF pump under run
UF pump Over run
No Flow
Low Flow
High Flow
Acid pump overrun
Acid pump underrun
Bicarb pump overrun
Bicarb pump underrun
High dialyzer inlet pressure
High dialyzer outlet pressure
Heater Temperature High
Communication lost
Low Arterial pressure
High Arterial pressure
Low venous pressure
High vinous pressure
Blood pump door open
Blood pump stopped
Blood not detected
Air bubble detected
Blood leak detected
Low dialysate temperature
High dialysate temperature
Low dialysate conductivity
High dialysate conductivity
Heparin pump stopped
High RO inlet temperature
BP underrun
BP overrun
Heparin underrun
Heparin overrun
Low RO inlet temperature
Issues